        2. General Settings for Full Closed Loop V.3-0 for3.0.1

Please note that with autoISF you are in an early-dev. environment, where the user interface is not optimized for safety of users who stray away from intended ways to use. Good safety features exist, but these are only as good as the development-oriented user understands and implements them. This is not a medical product, refer to disclaimer in section 0

2.1 SMB Range Extension
Available related case studies:
Case study 2.1: (nothing available yet)

2.2 Max and Min autoISF Ratio
2.3 SMB Delivery Ratio
2.4 iobTH
2.5 Eating Soon TT ?

When in Hybrid Closed Loop, big boli were given by the user, and the loop had no business to give big ones on top. But this is fundamentally different now.

So, first we must enable our loop to overcome the narrow safety restrictions for SMB sizes that were appropriate in HCL.
Evidently, this could become dangerous. Please set your iobTH (section 2.4), and observe all suggestions made in this FCL-e-book *) , and in the github pages of the developers.
See also: https://github.com/ga-zelle/autoISF/blob/A3.2.0.2_ai3.0/How-to-get-larger-SMBs.pdf

*) Caution: This entire e-book is about Full Closed Looping. In case you intend to work with giving boli, many suggestions made - notably in this section 2 (and in section 4) – should not be followed. You would have to do extra research, on your own data, how your bolus changes things. More see at end of section 4.1

2.1 SMB Range Extention
      (preferences/OpenAPS SMB/autoISF settings/smb delivery settings/smb_max_range_extension)

2.1.1 Standard scenario: SMB sizes based on 5 minute loop calculations

Full Closed Looping requires bigger SMB sizes. Setting 120 minutes of basal as max. SMB size as enabled in AAPS Master is necessary, but will rarely suffice.
In AAPS Preferences/OpenAPS SMB/autoISF settings/smb delivery settings, set SMB/UAM max range extention (smb_max_range_extension) to 2.0. That doubles the allowed max. size, to four hours of your profile basal, for a start.
Even better, you could determine an estimate for your initial setting as in the following description, I will use the symbol, ….. , to denote where you would use your numbers. My numbers that I use for the same situation will be in parentheses (U).
In full closed loop, once the bg starts rising, you want to get at least half of your required meal bolus within 10 minutes, through 2 SMBs. To do that, you need ..... U (2 U) per SMB on average, and because the bolus sizes tend not to be equal when requested by the Loop, you should have at least ..... U (3 U) as the allowable SMB size. Your hourly basal is around ..... U (0.6 U), i.e. AAPS Master will allow a max. 2 times that hourly basal which = ..... U (1.2 U) per SMB. To reach the intended .... U (3 U) therefore you should set your smb_max_range_extention to .... ( 2.5 = 3 U / 1.2 U)

The profile helper in section 4.7 might be available for doing this calculation, and for a cross-check.

If you have a very low hourly basal rate, extensions bigger than 3.0 can result.

Note: Elevated insulin needs in phases of elevated insulin resistance probably will be managed with >100% profile adjustments. Then, profile basal gets elevated accordingly, and thus will automatically allow increased SMB size.

Watch out (in your SMB tab, or using the emulator, section 10) whether you often run into a limitation by your set smb_max_range_extention. For instance, your attempts to increase initial SMB sizes via elevated smb_delivery_ratio (section 2.3) and elevated bgAccel_ISF_weight (section 4.2) might get cut by a too low smb_max_range_extention.

The SMBs your loop requests could get reduced in size also by other safety settings, notably by your autoISF_max setting (see section 2.2)

2.1.2 Special Libre 3 (1 minute) scenario with up to 5 SMBs per 5 minutes

When receiving bg values every minute, and adjusting insulin delivery accordingly in smaller steps, probably 120 minutes of basal per SMB suffices.
Leave SMB_range_extension at default 1.0 untouched. Can a Libre3 user confirm this, please ?
Integration of 1-minute values is new from autoISF 3.0.1 (May 2024). Please watch your SMB tab (difficult in 1 minute segments => make screenshots, or employ Emulator!), and report experiences / things we should include in the instructions.

2.2 Max and Min autoISF Ratio
        (preferences/OpenAPS SMB/autoISF settings/autoISF_min and autoISF_max)

For a start, set autoISF_max = 2.0 . –
Your CGM (1 or 5 minute re-calculations) should not matter for this parameter.
This allows up to doubling of ISF aggressiveness if "requested by the “… _ISF_weights" (see section 4). This is just a first step.
You may have to elevate autoISF_max further, later, if your attempts to tune the …ISF_weights (section 4.) often run into a limitation by your set autoISF_max

For situations of increased insulin sensitivity (less insulin need), you must specify, in your settings (preferences) also what your lower limit (for weakening of ISF, compared to profile_ISF) may be.

autoISF_min should be set to 0.3. Again, see whether you ever run into that limit, that your loop e.g. “would like” to act softer, but bounces against that set limit.

Do not keep autoISF_min near 1.0 because that would for instance pre-clude later, that your exercise setting can strongly “soften” your loop.

2.3 SMB Delivery Ratio
        (preferences/OpenAPS SMB/autoISF settings/smb delivery settings/smb_delivery_ratio)

2.3.1 Standard scenario, using 5 minute loop calculations

Use the fixed smb_delivery_ratio and increase the setting, from AAPS default 0.5, to 0.6.
This results in the algorithm generally demanding 20% more insulin (in the calculation it does every 5 minutes: insulinRequired * smb_delivery_ratio = requested extra iob).
Setting it to something like 0.6 is just a first step. You may have to elevate the smb_delivery_ratio further. For instance, setting it to 0.75 means, to get 50% more (compared to AAPS default setting) insulin regarding the portion of calculated insulin required that is delivered in the next SMB.
It is not recommended to go over 0.8 ever. The jumpier your CGM, the closer remain near 0.5!

In AAPS Preferences, the smb_delivery_ratio can also be set dynamically (changing with glucose level). - This seems a feature geared more towards hybrid closed loop applications. (In FCL, we like strong loop aggressiveness at low (but rising) bg, not “wait for” high bg).

2.3.2 Special Libre 3 scenario, using 1 minute loop calculations

1 minute increments lead, on average, to much smaller insulinRequired. As per step the ratio of random scatter to “true effect” may be more questionable, please start using a value around 0.2, and never go higher than 0.5 with your set smb_delivery_ratio, when using the 1-minute method.
(insert link to section 3 or to github libre3 document by ga-zelle) “

2.4 Safety Against too Aggressive Settings: iobTH
        (preferences/OpenAPS SMB/ autoISF settings/smb delivery settings/iob_threshold_percent
 …which gets multiplied with preferences/OpenAPS SMB/Maximum total IOB OpenAPS can’t go over (U)

A safety net is needed because autoISF shoots big SMBs when glucose levels begin to rise; but you do not want to bounce into your ultimate iobMAX safety setting too often.

Similarly, this safety net is needed also if using autoISF in Hybrid Closed Loop (HCL), where, after a user bolus already provided some iob, autoISFcould add too big SMBs to be safe.

Therefore, we install an iob threshold (iobTH) which, when, and as long as, exceeded, shuts SMBs off.

Step 1: In Preferences, set the SMB toggle for even /odd targets to “ON”

autoISF 3.0.1 demands (as we recommend for FCL also for other reasons, see e.g. later in section 5.1.2 and 5.1.3) that concurrently, in AAPS / Preferences / Open APS SMB / autoISF settings / SMB delivery settings:

        Enable alternative activation of SMB depending on current target *) ON

*) previous autoISF versions allowed different settings, now it is same setting, for profile target and for TT

Step 2: In Preferences, set your default iobTH_percent

iobTH is a parameter in AAPS preferences, defined there as fraction of your set maxIOB:
/OpenAPS_SMB/autoISF_settings/Full_Loop_settings: Percentage of maxIOB above which SMBs are disabled (iob_threshold_percent,)
Step 2.1: Solidify your maxIOB
First, check whether your maxIOB is set reasonably in AAPS Preferences / OpenAPS SMB / Maximum total IOB OpenAPS can’t go over (U).
Input a figure (units) slightly above the max level of iob you had ever needed in your past looping history (also considering times of elevated insulin resistance you occasionally may have had to deal with); set maxIOB to that value. A number around one fourth of your average TDD would seem plausible.

Step 2.2: Identify your max iob need in big meals
Now look at your meal spectrum, and what bolus size, and iob level (including from SMBs; in HCL or FCL) was useful *) in high carb meals to control your glucose. (For instance, the author needed up to 8U early-on in big meals in HCL; and he has TDD near 40 U, and maxIOB set to 10 U).
*) useful level = iob needed for the meal; iob may in time have gotten even higher. However, if, in the end, to prevent a hypoglycemia, you had to consume 15 g carbs, then deduct 15 g / (your IC) from that even higher iob you actually did temp. have. Example: 15 g / (10 g/U) = 1.5 U

Step 2.3: Set your iob_threshold_% in AAPS/Preferences
Assume you want to approach no more than about 75% of that iob level (that would be useful to have at big high carb meals) via rapid SMB „fire“, after a meal related bg rise is seen (then, for instance, reduce from 8 U to 6 U).

Then calculate your setting for iob_threshold_percent in AAPS / Preferences:
        = desired total iob given via SMBs before bg peaks / iobMAX

Enter the according percentage in /Preferences
(In the example it would be = 6 U / 10 U = 0.6 ; which means to enter 60 as percentage in /preferences ).

In section 4.7 a profile helper might be available for more guidance.

The iobTH then is calculated as follows:
        iobTH = iobTH_percent x maxIOB

Step 2.4: High-carbers may need to reduce their iob_threshold_percent a bit to factor in that the last “allowed” SMB can shoot above iobTH:
The last SMB given can exceed that threshold by up to +30% of the effective iobTH.
This is desirable because it allows higher iob at big high carb meals (where SMB size, when approaching iobTH, is still big); at lower carb meals either iobTH will not be reached anyways, or SMBs are quite small when reaching iobTH and will not shoot over by much.
A big SMB that would shoot over by more than +30% will be cut at 130% iobTH.
- Until iob falls below effective iobTH, only %TBRs supply more insulin, if the loop calculates that more iob is still required.
- In low carb meals, that iobTH level should not be reached => the autoISF parameters („weights“) need to be tuned carefully, so SMB sizes are not always huge, and bounce against the iobTH restriction, but show different behavior for different meals
- Note that when operating with an even elevated bg target (>100 mg/dl), iobTH can only be exceeded by +20% (“loop at medium power”). This makes sense, notably in an exercise context (…in which the iobTH per se also gets automatically lowered, as later discussed in section 6.1.3.).

Step 2.5: At some later stage, come back to fine-adjust your percent setting in /Preferences,, factoring in that it can be auto-modulated.
- autoISF 3.0 and higher contains a function to auto-adjust iobTH with TT set: Dynamic iobTH. This means, in your initial tuning, just set a iobTH_percent that is good-enough on your average day.
- The formula for the resulting effective iobTH is complicated, and also depends on whether the exercise mode (or the activity monitor) is active. It can be used to (semi-)automatically lower (or, in case of in-activity, elevate) iobTH (and hence iob) as highly desirable for exercise (resp. in-activity). See sections 3.3 and 6.1.3, and example in case study 6.2

- The resulting effective iobTH can be seen in the SMB tab.
  From autoISF 3.0.1 onwards, the SMB tab starts with the Result section, and right below the / start autoISF 3.0.1 / headline, the resulting modulated iobTH is given:

  In later software updates, it is desirable to see the modulated number (6.36 U in the example) also next to iob (below the glucose value in the AAPS main screen).

Green texts describe currently not available features that were suggested for further development

Likewise, a %profile setting < 100% (or > 100%) will proportionally lower (or elevate) the iobTH that will be used.
        effective iobTH = % temp.profile x iobTH
After the temporary % profile expired, it will automatically revert to your originally set iob_threshiold_percent.
Off-topic note, regarding the effective ISF (“sens”):
In the SMB tab, above the “start autoISF..” line, the profile ISF is given (“ISF unchanged”), eventually with adaptation by activity monitor (“adjusting …ISF from … to .. “ ? ) or by a TT (“adjusting …ISF from … to ..” ) or by a %temp. profile set (“unfortunately” still called “ISF unchanged” then).
Then follows the autoISF section explaining in detail how the recently encountered bg curve characteristics suggest adaptations, and what overall the conclusion is (“final ISF factor”, calculated following the flowcharts as explained in detail in section 03.).
Below the autoISF section, the effective ISF (sens) results from dividing the (unchanged or adapted) ISF prior to “start autoISF”, with the determined “final ISF factor” at the end of the autoISF section of the SMB tab.
Example given in section 5.4.5

If both, a % profile and a TT are set, both effects multiply.
We shall see later, how this opens nice avenues for exercise management, where we like to strongly limit how high iob shall be allowed to go. Example given in Case study 6.3.

All above discussed settings must be made in AAPS/preferences
except for the temporary modulations of sensitivity, which can be done from the AAPS home screen via %profile or via TT inputs
(This will in detail be presented in section 5 or, regarding exercise button, in section 6.1.3). Regarding activity monitor see section 6.6).

Setting different iobTH via an Automation.

If you do have situations where you rather use a different iobTH_percent than set in your profile (or than resulting in an eventual auto-modulation), you can change it also via an Automation.
Caution: This will overwrite your iob_threshold_percent until you restore it, manually (in /Preferences), or via another Automation (see e.g. section 5.1.4).
For this reason, preferably work with the temporary adaptions via %profile, TT and exercise mode as discussed above.

Next steps:

Before going first time into into FCL mode, you must first make additional settings in /preferences/Open APS SMB/autoISF, as discussed in detail in section 4.

Only after you also went through section 4, you can enter/exit FCL (for initial tuning, or for everyday utilization) via
- in AAPS/Preferences/OpenAPS SMB/autoISF settings/”Enable ISF adaptation by glucose behavior” ON / OFF
- or, in a suggested User Interface upgrade: by just tapping on the violet/ green closed loop icon of your AAPS home screen, after this feature (section 5.3.1) becomes integrated in a future autoISF version update.

2.5 EatingSoon TT ?

FCL works in principle also without setting an EatingSoonTT. Try for yourself whether you miss any performance via the totally hands-off way (and maybe skip this section for now).

Your FCL works best if you start meals at below-target glucose values, and ideally have a bit of positive iob at meal start. Also, a low temp. glucose target helps making SMBs (that „aim at it“) a bit bigger.

Setting an EatingSoonTT well ahead of meal start therefore is in principle a good idea

- If you have relatively fixed meal time slots in the 24 hours of the day, you could set the target glucose values in your profile accordingly. So e.g. 11-15h target 76 mg/dl if you almost always start a lunch between 11:45 and 14:30h. (If you do exercise or physical work in that time, this would be too aggressive, and probably also un-necessary).
- If you have rather irregular habits, it might be worthwhile to manually set an EatingSoonTT (which is quite time-uncritical) well before the start of a meal, or even (latest) when the first SMB is about to be triggered by your loop. – Or, just forget about it:
However, if (as to be expected) your loop anyways always regulates you down to near-target, the effect from setting an EatingSoonTT will be limited:
- If your meals are spaced by a couple of hours your glucose should not be elevated as you approach the next meal …
- If you eat more at a still elevated glucose, your loop should provide you with “a balancing” iob (and hence a prediction to get to target soon) …
…In both cases (which are the prevalent norm) setting an EatingSoonTT would only have a very minor effect.

The good news therefore is: Setting an EatingSoonTT has only minor effects, if any, and is not required for autoISF FCL.
Preferred solution with respect to Eating SoonTT

My preferred solution is as follows

Principal solution: I am not setting an EatingSoonTT ahead of the meal. For getting full loop aggressiveness, an even profile target (or TT) under 100 mg/dl is the only requirement.

Note: To prevent full loop aggressiveness when not needed, e.g. at night time, I set an odd profile target (or TT); details see section 5.1.2.

Setting odd profile (or temp.) target is generally a recommended easy way for you to block out SMBs, as an “emergency brake” for your FCL. This should be particularly useful in your initial tuning work.

Refinement: Just “to orient the calculated insulinRequired towards a more aggressive target”, I have my loop automatically set a low TT just around the time when it gives first SMBs

For this you can define an Automation like: CONDITIONS: likely meal time of day + 1st sign of a beginning meal + iob under (?) U (to kick in only in initial phase) + no TT running => ACTION: set TT=74 mg/dl for ~30 minutes

This Automation does not make much of a difference, really. You could leave eventual implementation of this “refinement” to a later round of fine-tuning things.