    4. Meals: Setting ISF_weights in /Preferences v.3.4 for 3.0.1

        V.2.0

Please note that with autoISF you are in an early-dev. environment, where the user interface is not optimized for safety of users who stray away from intended ways to use. Good safety features exist, but these are only as good as the development-oriented user understands and implements them. This is not a medical product, refer to disclaimer in section 0

Available related case studies:

Case study 4.1: Pizza
Case study 4.3: Hands-off FCL on Xmas

  4.1 Getting started
  4.2 Initial bg rise: bgAccel_ISF
  4.3 Strong bg rise: pp_ISF
  4.4 Sluggish bg rise: bgBrake_ISF
  4.5 High bg: dura_ISF and bg_ISF
      4.5.1 dura_ISF
      4.5.2 bg_ISF
     4.5.3 How “UAM” concludes insulinRequ.
     4.5.4 Managing high bg
  4.6 Tuning your initial settings
  4.7 Covering more complex scenarios
  4.8 Profile helper

Warning regarding importance of proper profile ISFs.

Starters on autoISF FCL who are coming from using HCL with dynamicISF must be aware of the following: It is absolutely essential to build your FCL on properly set profile ISFs (likely a circadian pattern over 24 hrs).
It may not apply to you, but many dynamicISF users did never bother to determine their ISFs that would maximize their HCL performance, but employ dynamicISF so to speak for going „dynamically“ through a wide range of possible ISFs, until eventually hitting a sweet spot, and the whole thing works better than before, with what they had used as a profile ISF (often only one, e.g. coming from Autotune).

The following is important to understand, as it leads straight into the core idea behind FCL with autoISF, too: It is a good idea to establish a well-running hybrid closed loop with set (non-dynamic) ISF (set in profile for each hour of the day). That ISF must be aggressive enough that it gets you down from a high around 200 mg/dl to target. That is roughly also the way you experimentally determined it (I hope. See https://github.com/bernie4375/HCL-Meal-Mgt.-ISF-and-IC-settings/blob/HCL-.-settings-main-repo-(pdf)/ISF%20determination_V.3.33.pdf ).
- Using that strong value also at lower bg, (on the way “up” , after meal start), is very positive: We do not want to have a softer acting loop when at lower bg (which is what dynamicISF tends to do!). autoISF will, in contrast, temporarily sharpen your ISF when, at low bg, acceleration is detected..
- On the way down from peak, towards glucose target, a somewhat too strong ISF should not hurt because much of the time your loop (well supplied with insulin before, „on the way up“) is zero temping , or at least has only a small gap to correct, from predicted bg to target bg.
- You have no business to be much above 200 mg/dl where an even stronger ISF may or may not help. It sure does not help at an occlusion which is about the only reason to see super high values as an experienced looper.

Pegging ISF strength to bg level therefore does not make sense in FCL. You will use the autoISF toolbox to get strongest ISF at low but beginning-to-rise bg,

Note: There are very much refined versions of dynamicISF that can have beneficial applications, notably in HCL.

Going to autoISF FCL, you absolutely must anchor on the proper profile_ISF.

When using autoISF you can – as you did in the past, e.g. around exercise, or in times of illness – temporarily modify your profile ISFs, via a %profile switch. Also the other two top buttons, exercise and TT, can be used to adapt to changes in sensitivity/resistance. More about that in section 5.2.2.2 . But, first, spend a couple of days (if not weeks) to get your key autoISF related settings right, strictly on/for days with your normal insulin sensitivity. This is what this section 4 is about.

Warning: Do not copy settings from other FCL loopers

When setting your parameters, don't use any given numerical example (not even as “a starting point”). Instead, anchor on data from your successful Hybrid Closed Loop!

Most examples given in this paper are from an adult diabetic (Lyumjev, G6) whose insulin sensitivity can be characterized as follows: approximately 37 U TDD, thereof 13 U profile basal, at about 200g daily carbs from mainly lunch and dinner; no couch snacks or sweet drinks. The user also participates in multiple instances of daily moderate exercise such as dog walking, biking and gardening. In Hybrid Closed Loop, a typical meal bolus was 8 U that was sometimes reduced such as when activity followed the meal.

After seeing some more inputs from a variety of users we might put together a profile helper for some rough orientation, and for plausibility cross-checking, in section 4.8

Warning. Importance of starting from a well-performing Hybrid Closed Loop

A satisfying performance in Hybrid Closed Loop mode is a pre-requisite. Expect to reproduce about the same %TIR also in your FCL, but with less daily interaction, once established.
Note that this refers to prior use of „vanilla“ software, without fancy „dynamic add-ons“ (such as: Autotune determined factors, dynamicISF etc). that may have introduced bias into the profile settings you bring with you into FCL now.

To reach a satisfying performance you must start from a hybrid closed loop in which you did master your meal management well using the oref(1) algo SMB+UAM. This is a pre-requisite to be able to forget it … - because the initial tuning that we now turn to demands, that you analyze your prior best practice, in an attempt to find appropriate settings and „teach“ your FCL to come up with the necessary iob.

This is the main subject of this section 4 (finding settings for automatic meal management).

 Section 5 will explore avenues to manage “disturbances”, i.e. time blocks or situations that might demand enhanced or reduced loop aggressiveness.
Section 6 will focus on the exercise mode, and the activity monitor.

Resist the temptation to make use of the tools presented in sections 5 and 6 too early.
On your first setting-up and tuning attempt, it is strongly recommended that you not “play around” with all ultimately available features, but stick to the sequence of steps to take.

Yes, “playing around” with the many extra buttons often will help find an improvement. But you likely create an instable FCL that, already at fairly standard situations, uses up some of your FCL’s principal capacity to correct for disturbances. This limits what will be left to manage extreme situations.

Also, once you created a maze of little errors and counter-strategies/counter-errors, it will be nearly impossible to find your way out of this mess, towards better settings, at any later point of time.
AutoISF comes with very many extra parameters, and even when employing the Emulator (sections 10 and 11) it is quite hard to analyze their interaction.
One principal reason why things are difficult to analyze is, that you really can only analyze one change, and that will put you on another bg curve. So, you can never see the full effect, along more than 10 minutes, that any change will ultimately result in.

PS: Section 11 describes the ultimate tool to investigate “what-if” regarding a setting change you may contemplate.

        4.1 Getting started

Make sure you have studied the preceding sections 1-2 on the general pre-requisites for FCL and section 3 on the principal workings of autoISF.

Caution: This entire e-book is about Full Closed Looping. In case you intend to work with giving boli, many suggestions made - notably in this section 4, and in section 2 – should not be followed. You would have to do extra research, on your own data, how your bolus changes things.
 (See also section 07, and discussion on pre-bolussing, ~2 pages down)

Make sure you have appropriately:

- widened the SMB size restrictions (section 2.1),
- elevated the max allowed ISF amplification with your set autoISFmax ( section 2.2)
- set your iobTH% (refer to section 2.4 and if available 4.8)

In the early test phase, it is recommended to:
Run the system as dummy, not connected to your body (or, on own risk, connect only as long as you watch closely)
In AAPS preferences, switch your autoISF FCL ( = autoISF/”Enable adaptation of ISF to glucose behavior”) ON only during daytime hours of a meal, e.g. 11-18h, for fully automatic "full closed loop" management of lunches.
You can do this switching manually at 11 h and 18 h every day, or set up an Automation that does that (see section 3.4 ).
- Take typical but not extreme meals. Omit sweet drinks, or drink only slowly.
- Occasionally, watch the time-pattern of bg, iob (SMBs given), and insulin activity after meal start. Aside from serious “mathematical” attempts to tune settings based on data from the SMB tab (or the emulator, section 10), just watching the curves develop on your AAPS main screen can, over time, give you “a feel” what settings, and eating behaviors, are benign or detrimental to good %TIR performance.
- It is wasted time to “optimize” settings based on 1 type of meal. You need a “good enough” compromise that works with your range of usual meals. See case study 8.2
- Do not use the Activity monitor (see section 6.6), unless it is already well calibrated. In case you use an EatingSoonTT at meal start, note that any active TT shuts activity monitor automatically off for a while.

The core challenge of your UAM Full Closed Loop is to recognize a meal start from the glucose trend, and ramping up iob.
When setting up your autoISF Full Closed Loop, you must set several ISF_weight parameters in AAPS Preferences/OpenAPS SMB/autoISF settings.

They relate to different stages of the typical glucose curve after starting a meal:

Note: bg_ISF is not used much in FCL, as it is rather late to act on high (or low) bg level that developed. But, feel free to experiment, e.g. in case you have indications, in your data, that in the past dynamicISF was useful to manage bg extremes in some situations.

The core advantage of using autoISF with oref(1) SMB+UAM (in FCL as well as in hybrid closed loop) is that it manages the glucose curve it sees developing, no matter what the underlying reason is.
42 potential factors were identified (see: https://github.com/bernie4375/HCL-Meal-Mgt.-ISF-and-IC-settings/blob/FCL-w/autoISF/42%20factors%20influence%20bg.pdf ), so, no wonder, that loopers who meticulously input their carbs will often not see the expected result.

Reminder: autoISF has that advantage only if the pre-requisites (section 1) are given, notably a very fast insulin, and reliable CGM and insulin delivery (not leaking, and permanently Bluetooth connected).

Before you progress, make sure you studied the flowcharts in section 3 that describe how autoISF calculates the effective(ly used) ISF.
Consult this also to help you understand how, in your SMB tab, the applied effective ISF is calculated.

Warning: Any bolus you „sneak in“ will severely distort the glucose curve. That can render your tuning of weights (see below) useless, and could make your loop act in unpredictable and potentially dangerous ways:
In case you feel tempted to use boli, be ready for some own extra research, read the inserted grey box, and also refer to section 7)

- Your FCL ideally runs without an insulin button at the bottom of the AAPS home screen.
- Issuing a bolus should kick you out of the FCL mode, back into Hybrid Closed Loop. We are working on improving the User Interface (see e.g. section 5.3.1) that would facilitate and secure the bi-directional transitions.
- How proper autoISF settings would differ for your meal management, depending on
  - no-bolus (FCL),
  - bolus like in HCL,
  - or very small pre-bolus (Meal Announcement),
is not well investigated at this point.
Therefore it is best to remain consistent = use autoISF strictly for no-bolus FCL, and if you want to bolus for a meal, switch ISF adaptation to glucose behavior (autoISF) temporarily OFF.
Maybe we are too cautious here, and in fact the autoISF adaptation to glucose behavior is tolerant enough of disturbances by user boli. Please report your findings in case you collect data of “mixed use” (FCL / Meal Announcement / HCL use with meal bolus).
A n=1 finding, and guide how to evaluate, is reported here: https://github.com/ga-zelle/autoISF/blob/A3.2.0.2_ai3.0/To%20prebolus%20or%20not%20to%20prebolus.pdf ).
Once we have a body of data, including from those who moved from HCL with autoISF to FCL, we may need to re-define what the bi-directional transitions FCL < - > HCL in detail shall mean, and whether or not this has implications for needing different autoISF settings in /preferences for FCL and for HCL. .

After doing the prep work as outlined in section 2 you now get to calibrate your FCL to your normal meal spectrum by initially setting and tuning the various _ISF_weights, that dynamically change with bg curve characteristics as sketched in the chart on the previous page.

Please stay away from extremes (regarding both, meals and exercise) when you go through this section 4. It is about getting a first roughly right set of settings, as a basis.

Researching your standard meal patterns, and finding settings for the various -ISF_weights is the core job in setting up your autoISF FCL.
Depending how varied your diet and general lifestyle are (and your expectation of %TIR you like to reach), this could be the main job at hand. However, there is much more you could do later, and that will be outlined in later sections 5 and 6.

        4.2 Meal detection and managing the initial bg rise: bgAcceI_ISF

When looping without carb inputs and without giving a bolus ourselves, the first crucial setting is to set the bgAccel_ISF_weight so that large SMBs are requested immediately when the loop detects an acceleration in your blood glucose (bg) that is starting to rise.

Ideally within about 20 minutes after acceleration detection, which would be the first up to 4 SMBs, as much iob should automatically be supplied as we would have given with our bolus in hybrid closed loop.

Rule of thumb: Two of the first three SMBs each should be about ¼ to 1/3 the size of a previous big meal bolus in your HCL „career“.
Going over 1/3 would be problematic if your diet contains occasional low carb (or only snacking), and generally of course if your CGM quality is sometimes unreliable, and might produce an artefact that could be mistaken for a meal start. Be vigilant about this topic!

For hands-off FCL, your settings have to fit the whole range of your meals in each of your meal times, e.g. should suit (nearly) all your lunches that you tend to have.
Between your daily mealtime slots, your circadian ISFs make a differentiation.
In extreme cases you will have to balance too high running iob with additional carbs (a late additional snack against going too low), and in the opposite case, you will have to reckon with temporarily exceeding the glucose target range, and losing some %TIR for this day.

If your meals vary very strongly, there are avenues to ease your initial tuning job, or to optimize overall resulting loop performance:
Automations allow you to differentiate. For instance it is possible to apply different iobTH_percent and/or different bgAccel_ISF_weights for meals in different time windows or geo locations (details see sections 3.4 and 5.1).
In case you use autoISF on the iAPS/Open iAPS platform for i-phones, you may need to use a third party automation software, or “middleware“ (! call for a case study 4.X )
You can pre-program custom buttons for special meal (or snack) types, with different underlying FCL settings (see “cockpit”, section 5.2.2.3)
Green texts describe currently not available features that were suggested for further development

In an update, autoISF 3.x might provide the option to pre-program settings for 4 different meal type clusters, accessible from the TT button (presented in section 5.3.3.1 (4) and 6.3).
You can modulate FCL aggressiveness manually making use of temporary switches of %profile and/or set for a a couple of minutes an odd (=>SMB off) glucose target (section 5.2.2.2)
Experimenting with the three above mentioned “avenues”, the author found:
- the third easiest to occasionally use, and the first one hardest.
- it worth investing some effort (also using the emulator a couple of times) to iterate through the typical meal spectrum a couple of times, for finding a “good enough” set of .._ISF_weights and other settings (like autoISFmax, iobTH% etc), and not do much extra differentiation. (More see in section 5).
It is certainly worth trying hard at finding a good set of ISF_weights for your meal spectrum, to keep interventions in daily life to a minimum.

In search of appropriate settings, you must keep (real-time) track of the SMB tab when tuning. This can be impractical. You probably will end up making a lot of screenshots (quickly in the crucial minutes after a SMB was given, or when you thought it should be given), for later analysis.

The superior method is to just copy logfiles
...about once a day, or from autoISF 3.0.1 onwards, every few days suffices...
from your phone/internal memory/AAPS/logs (all zip files there), and analyze them at your convenience later, using the emulator (see section 10; used e.g. in last pages of case study 4.1). Some emulator-based analysis is also possible within AAPS on your phone (section-11).

Already when tuning the bgAccel_ISF_weight it can become evident that safety restrictions (as discussed in section 2) must be widened further:
- Especially if your profile basal rate is very small, the smb_max_range_extention and/or the autoISF_max "must" often be increased further.
- Furthermore, the smb_delivery_ratio provides more leeway to increase the aggressiveness (e.g. 0.6 -> 0.72 results in another +20%).
In the end you should not set these safety limits too tight, so “nudging” aggressiveness by another 10 or 20% from your cockpit, later, will not bounce into restrictions.

On the other hand, setting narrower restrictions for max allowed SMB size can also become necessary:
- You don’t want your loop bounce, regardless of the carb load, “immediately” into your iobTH limit (and up to 30% above), which is not desirable if your meal spectrum is very varied
- Poorer CGM quality demands narrower restrictions, too, for safety reasons.
- If you use a 1-minute CGM (Libre 3) please observe section 1.4.2

In any case, it is worth the effort to tune the bgAccel_ISF_weight in such a way that high glucose increases are already nipped in the bud, so to speak.
Remember: In FCL, the first 3 or 4 SMBs should not be much delayed, and amount to similar iob like your “former boli in HCL”.

Early strong iob also eases the tuning task for the subsequent phases of the meal, because there is, then, largely zero-temping, as well known from HCL-times after your administered bolus. Also, the lower and shorter lasting the glucose peak, the lesser the hypo danger from the activity tail of SMBs given when glucose was „stuck“ high.

bgAccel_ISF_weight is set default to zero in autoISF.
To start, I would try 0.05 or max 0.1, and keep trying in max 0.05 steps. Soon move to 0.02 steps. From my (very limited) overview, many use around 0.2, and possibly even higher if their hourly basal rate is 0.1U or lower. (Consult section 4.8 when available). Do not be tempted to rush this setting by using large jumps in adjustments.
How changing the _weights influences the resulting calculated insulinRequired

To get a feel for how changing the _weights influences the resulting calculated insulinRequired, it is best to start cautiously and just do 10 to max 20% steps up, and watch out for the effects. Doing similar step sizes should yield about similar effects each time.
Example 1: Going from bgAccel_ISF_weight of 0.2 to 0.16 (20% less).
If your profile_ISF is 40 mg/dl/U and with bgAccel_ISF_weight = 0.20 you saw acce_ISF factor of 1.31, this would ((if the acce influence dominates and is used as effective ISF)) lead to the effectively used ISF of 40/1.31 = 30.53 mg/dl/U. For an intended correction by – 10 mg/dl the insulinRequired would calculate to 10 / 30.53 = 0.328 U.
Now, going with a 20% reduced bgAccel_ISF_weight of 0.16:
        acce_ISF = 1+ bgAccel_ISF_weight * internalFactor
before 1,31 = 1 + 0.20 * iF => 0.31 = 0.20 * iF => iF = 1,55
after ? = 1 + 0.16 * iF => ? = 1 + 0.16 * 1.55 = 1.25
New effective ISF would be 40 / 1.25 = 32.05 mg/dl/U. For an intended correction by – 10 mg/dl the insulinRequired would calculate to 10 / 32.05 = 0.312 U, which is 4.9% less.

Example 2: Going from bgAccel_ISF_weight of 0.2 to 0.10 (50% less; or doubling in the other direction).
If your profile_ISF is 40 mg/dl/U and with bgAccel_ISF_weight = 0.20 you saw acce_ISF factor of 1.31, this would ((if the acce influence dominates and is used as effective ISF)) lead to the effectively used ISF of 40/1.31 = 30.53 mg/dl/U. For an intended correction by – 10 mg/dl the insulinRequired would calculate to 10 / 30.53 = 0.328 U.
Now, going with a 50% reduced bgAccel_ISF_weight of 0.10:
        acce_ISF = 1+ bgAccel_ISF_weight * internalFactor
before 1,31 = 1 + 0.20 * iF => 0.31 = 0.20 * iF => iF = 1,55
after ? = 1 + 0.10 * iF => ? = 1 + 0.10 * 1.55 = 1.155
New effective ISF would be 40 / 1.155 = 34.63 mg/dl/U. For an intended correction by – 10 mg/dl the insulinRequired would calculate to 10 / 34.63 = 0.289 U, which is 12 % less (going the other way, 0.328 is 13.5 % more).

Example 2 (-50%) reduces _weight 2.5 times lower than example 1 (-20%), and the resulting effect (-12% vs. -4.9% insulin Required) is also factor 2.5 different.

Note: “Your” internal factor “iF” might differ; for sure it is very different between the various …_ISF components. (Also, never forget to look into how other .._ISFs play into the effective_ISF which overall results).

Ideally, one should set the bgAccel_ISF_weight such, that for meals that are in the lower (!) range of the "fast carb load" of your cluster, the necessary insulin supply is already approximately provided with 3 SMBs.

The glucose curve, at such meals, begins to flatten early in this SMB phase, so a de-celeration (braking) follows very soon (-> section 4.4).

Note regarding acceleration happening “again” in late part of dropping glucose

After the peak, in the late stage of falling bg, the glucose curve is like an accelerating parabola again. The algorithm tries to evaluate when and at which bg level complete digestion of the meal and a bg minimum will result. Insulin required to stabilize around target bg is usually very small, and the adaptation of ISF in that stage relatively unimportant.

4.3 Managing strong bg rises: pp_ISF

With higher carb load meals, or meals that come with a sweet drink, the acceleration phase will last longer, and bg will rise further, which will require a higher insulin supply.

Between acceleration and deceleration there is a more or less linear further increase of insulin need in these cases.

autoISF should now "fight" this with the help of the post-prandial ISF, set via pp_ISF_weight, after we have set a halfway suitable bgAccel_ISF_weight.

Tune your pp_ISF_weight after you have set a halfway suitable (not too aggressive) bgAccel_ISF_weight. You now should check meals in the upper spectrum of your g carb load, and carefully start with a starting value for pp_ISF_weight of 0.005. Observe the reactions and check the SMB-tab before you increase it with care for the next day.s
Best practice is to analyze the emulator tables (discussed in section 10, and example given in the pizza case study 4.1)

Normally (except for very low carb meals) the SMBs triggered by bgAccel_ISF_weight and pp_ISF_weight should be sufficient to reach and slightly exceed the iobTH (see section 2.4) so all the other autoISF parameters are relatively unimportant for now.

A reason why this can work at all, also for quite a variety of meals, lies in the fact that there is an hourly carb absorption limit of about 30g/h
(Reference: Dana Lewis:https://github.com/danamlewis/artificialpancreasbook/blob/master/8.-tips-and-tricks-for-real-life-with-an-aps.md#heres-the-detailed-explanation-of-what-we-learned. (That limit can be lower, e.g. with gastroparesis or certain medications, but that would make things even easier)

So while meals might wildly vary in composition and size: What is digested, and needs insulin in the first ~90 minutes (when FCL tries to catch up with insulin need and differs strongly from HCL, with bgAccel_ISF and pp_ISF in the leading role), will be relatively close…for meals with similar initial glucose acceleration and rises, anyways…

The others, low carb with much slower initial acceleration and rise, are easy recognized as different by the loop, see section 4.4 that follows.

Depending on the type of meal and "aggressiveness" of your bgAccel_ISF_weight and pp_ISF_weight tuning, the iob will already be so high that, in the phase of decelerated glucose rise towards the peak (the "last part of the rise"), no more insulinReq is seen by the loop.

Therefore the bgBrake_ISF_weight is often unimportant in meals with a relevant carb content.
For relevance in low carb meals, see section 4.4.

Warning: Occasionally consult the SMB tab to see how your settings really work.

A setting (...ISF_weight) that is actually set too aggressive might be masked. Tuning only works if the effects of the settings being tuned are not unintentionally limited by other (e.g.„safety“) settings.

Also, always look at two or three different meals before deciding whether a tuning "fits" („good enough“ for each of them). You probably will have to iterate back and forth doing this for two or three different kinds of meals …

- Case Study 4.1 (Pizza Meal) contains, towards the end, an example how you can go about tuning the _weights for various _ISF factors of autoISF.
- Case Study 8.2 shows that it is not worth it to seek “optimized” settings based on just one meal.

… until you find one good enough set of settings for all of them. Do not rush this, establishing a solid foundation will be well worth your time.

       4.4 Sluggish rise towards bg peak: bgBrake_ISF

At a low carb meal, or an attempt at doing a weight reduction diet, (and probably also with gastroparesis, or if you take one of these novel GLP-1 drugs that slow meal absorption - Somebody, please supply a case study!) the glucose goes up only sluggishly, and iobTH should not be reached at all.
In case you always do very slow absorbing meals, you could of course also adjust your iobTH setting low enough to suit your uniform situation.

Acceleration, and the phase of strong glucose rise, are quickly over at slow-absorbing meals, and there is mainly a decelerating bulge of insulin action that projects over the next few hours.
Now the importance of the bgBrake_ISF_weight comes in.

In full closed loop, the bgBrake_ISF_weight is often only about half as large as the bgAccel_ISF_weight (but that would also depend on your personal diet pattern and eating/digestion speed). Also here, one should approach the tuning gradually, increasing the weight coming from small values.
Please observe that this tuning must strictly be done with types of meals for which there is insulin need at de-celerating but still rising bg.
bgBrake_ISF is totally irrelevant for hi carb meals where your loop shot over iobTH already by the time your rising towards the bg peak slows down!
Likewise, if your initial bgAccel_weight is set so strong that your first SMBs catapult you over the iobTH, no matter what type of meal: Then you must first find a reasonable setting for this parameter, one that works “good enough” to control your carb loaded meals, but still leaves room for milder loop response at low carb meals.

In case you cannot quite get all the ISF_weights “right” so the occasional low carb meal will not get over-treated: Avenues to adapt your loop aggressiveness are discussed in section 5. For indstance you will be able to (if needed);

- use a temp. reduced %profile
- temp. lower iobTH or bgAccel_ISF_weight
- construct for yourself an extra snack or low carb button (“DIY cockpit”) with an underlying suitable Automation

In the late stage of still rising (!) glucose, the Full Closed Loop typically sharply reduces SMBs already because it is “painfully aware” of the following principal conflict:

- iob (like formerly given in HCL via your bolus) must go high quickly, in order to limit the high
- However, if there is too much insulin in the system, a hypoglycemia can happen later within the DIA time window, because the loop can, later, only correct to a very limited extent (namely, only to the extent that it can set basal to zero).
Therefore, the core problem is that the Full Closed Loop must build up iob very quickly, but not too much, in the initial phase of a meal, and high bg values (out of range, >180 mg/dl) can not always be avoided.

4.5 High Glucose Values: dura_ISF and bg_ISF

4.5.1 dura_ISF

With large or high fat/protein meals, a 2nd hill will form in the bg curve, or a long high plateau.
For such situations, autoISF features the modulation of ISF depending on bg level and duration of plateau formation.

High plateaus in bg values are addressed via tuning the dura_ISF_weight and associated parameters.
This feature is also very useful in Hybrid Closed Loop. It can be used to elegantly manage, fully automatically, a temporary insulin resistance from fatty acids. Please refer to other papers for details (for instance, section „Late stage of meals“ of “Meal Management Basics”, available here: https://github.com/bernie4375/HCL-Meal-Mgt.-ISF-and-IC-settings ).

Set a start value of 0.2 for your dura_ISF_weight, and increase only cautiously with an eye on hypo prevention 2-3 hours later. Fine tuning this parameter only makes sense after you tuned your bgAccel_ISF well (so your thin yellow insulin activity curve shifts as far to the left, towards meal start, as possible, which will lower bg peaks and ease the job for dura_ISF).

4.5.2 bg_ISF

Since in Full Closed Loop we make our loop give us the maximum SMB size it can, at the beginning of a rise, it is crucial to resist the temptation to continue with a particularly strong ISF in the meal phase with the highest glucose values .
This is a reason why in Full Closed Loop we do not make much use of the bg_ISF component of autoISF.
Wanting to get most of our insulin from SMBs delivered at fairly low (but beginning-to-rise) bg implies that we do not make ISF weaker at low bg. Under preferences/OpenAPS SMB/autoISF/bg_ISF settings you could set lower ISF_range_weight = 0.0.
If you want to analyze in your data, whether you might benefit from a milder ISF at low bg values (e.g. if you often go below target after correction of only mildly elevated bg in the preceding hours), you may want to try lower ISF_range_weight = 0.1 or 0.2. Study the effects from bgISF, and increase, or decrease, the bgISF_weight to fine tuine the sought-after affect.
The higher_ISF_range_weight is used when bg is above target, It then strengthens ISF the more the higher the set weight is. 0 disables this contribution, i.e. ISF is constant in the whole range above target.
In FCL, this factor should be fairly irrelevant: Near glucose peak, zero-temping usually prevails anyway, so the settings we try might often not be used really by the loop. Very likely, you can live with setting the weight to = 0.0 here, too.
If you want to analyze in your data, whether you might benefit from a stronger ISF at high bg values (e.g. if you often remain above target after correction of elevated bg in the preceding hours), you may want to try higher ISF_range_weight = 0.1 or 0.2. Study the effects from bg_ISF, and increase, or decrease, the higher_ISF_range_weight to fine tune the sought-after affect.
Caution: Investigating effects of set weights is not really possible in periods of zero-temping. Too aggressive settings might not come into play most of the time. However, some other time they might come into play, and then produce a hypo 1-2 hours later.
Therefore, carefully study the SMB tab (or better yet, do an emulator based analysis, see sections 10-11) to see what the selected weights would do, if there was no zero-temping at the time.

Also, try a completely different meal, to see how your settings work there.

4.5.3 How your “UAM” concludes insulin need for your un-declared carbs

The UAM Full Closed Loop doesn't get any information from you as to how many grams of carbs will be there, for absorption. Not knowing when your steady-state max carb absorption phase
  - the earlier mentioned 30g/h, or
  - with gastroparesis, or if on GLP-1 drug treatment, probably on a lower g/h level
  - sometimes complemented/prolonged by a brief episode of insulin resistance to fats
might end, the FCL will struggle to provide desired amounts of insulin, facing potential hypo danger later (because of the DIA of the insulin in use).

Actually, the UAM Full Closed Loop is not completely clueless regarding how carb absorption will go on.
It will work with a prediction of further carb absorption building on the carb deviation (=hypothesis of how much got absorbed in the past 5 minute segments), and phase out further expected carb decay in the course of the next 1 to max 3 hours. For more detail see https://openaps.readthedocs.io/en/latest/docs/While%20You%20Wait%20For%20Gear/Understand-determine-basal.html#understanding-the-basic-logic-written-version (or study your SMB tab info).

This UAM prediction about further carb absorption can be worse, but can also be better than a prediction based on the user‘s „e-Carb“ input in Hybrid Closed Loop.

In any case, and even when having perfect knowledge about how exactly the carbs fade out in the next hours, there would still be a principal problem for the loop: Heavy insulin „fire“ against highs will not work immediately (depending on the insulin’s time-to-peak), and notably it comes with a significant hypo danger (from the „tail“ of insulin activity.)
A big bolus, or also a series of boli, will rarely work exactly for several hours matching the absorption of carbs (from what, how much and and how fast the user ate).

4.5.4 Conclusion on managing bg highs

Once your bg sits high, neither you, nor a hybrid closed loop with all the carb info, nor your FCL can work wonders.

Resist the temptation to elevate the dura_ISF_weight very high.

The author is sceptical about using the bg_ISF in Full Closed Loop:
- In FCL you probably can afford to shut it entirely off via setting both related _weights to 0.0.
- At least be careful, use small ISF_range_weights and check whether you are happy with the contributions to effectively used ISFs
  - Off topic: If you come from dynamicISF usage to HCL with autoISF, you probably can use the bg_ISF parameter with higher _weights to emulate what you like to replicate from your dynamicISF experience.

bg highs will take time to resolve.
Interestingly, an after-dinner walk can work wonders sometimes (take glucose tablets along).

   4.6 Tuning your initial settings

Be pro-active: The earlier large SMBs come (driven by bgAccel_ISF and pp_ISF) …
Note: Also your CGM smoothing may play a role here, that you may want to look into !
…the less high the overall increase in BG will be, and (provided you set a proper iobTH) the lesser the risk will be for a hypo after the meal.

Therefore, put most of your FCL tuning effort into determining suitable weights for bgAccel_ and for pp_ISF, and for finding a suitable iobTH_percent.

Later, your FCL cockpit will give you access to temporarily modulate these essential parameters (see section 5.2.), providing you an opportunity
- in your tuning phase, for more research on the fly, so to speak
- everyday, for temp. adaptations to altered insulin sensitivity, or to special disturbances (if you occasionally see a need).

After you tuned your initial settings well, there should rarely arise a need for “fine tuning” later, see section 8 and case study 8.2!

The experience of the author is that it is possible to tune the above mentioned weights for very different meals in such a way that the glucose almost always remains acceptably in range.

However, if you come to the conclusion that differentiated settings (for different meals or meal time clusters) would be easier to establish, and/or work better for you, the following sections suggest many options you could try and use.
4.7 Covering more complex scenarios

You now can move on, to accommodate more complex scenarios.

- Depending
  - how satisfied you are with your initially reached result, or which more extreme meals (smaller? faster/slower carbs? totally different fat/protein content?) you would like your FCL to manage as well, or
  - whether you seek temporary adjustments that make your FCL act more aggressive, or softer
you have a variety of options to deal with that, and this will be the topic in section 5.

It is suggested to do major exercise still in your hybrid closed loop setting, until you have your FCL up and running for meals on normal days with no or only moderate exercise. Later, implement extras as discussed in section 6 to fully implement your FCL.

To deal with different disturbances than presented by the meal spectrum you were calibrating for, there will be temporary modulations of your FCL possible.

Manual, making use of the top 3 buttons (%profile, exercise, TT; TT;section 5.2.2.2) or
- Semi-automatic (user triggered), aided by Automations you would set up, with a user defined extra button in your cockpit for it (section 5.2.2.3) or
- fully automatic (via pre-defined settings and/or Automations that e.g. that use different iobTH and/or different bgAccel_ISF-weights for different rough meal-time slots in your days: section 5.1.4 )
In future autoISF versions we could also pre-program 4 different clusters in /preferences, and call them up within a second from the TT button in the AAPS home screen (only after implementation of an improved cockpit, see section 5.3.3.1 (4) and section 6.4.3))

So, while FCL is about fully automatic cruising, your AAPS main screen will serve you as your cockpit to check how everything is running, and to aid your loop manouvering through some special disturbances.

In the SMB tab you can see how the autoISF modulation of ISF is overall applied to arrive at the actually used effective ISF (“sens”): See also example given in section 5.4.5

In the SMB tab, above the “start autoISF..” line, the profile ISF is given (“ISF unchanged”), eventually with adaptation by activity monitor (“adjusting …ISF from … to ..) or by a TT (“adjusting …ISF from … to ..” ) or by a %temp. profile set - (still called “ISF unchanged” then, meaning unchanged yet by autoISF).
- Then follows the autoISF section, explaining in detail how the recently encountered bg curve characteristics suggest adaptations, and what overall the conclusion is (“final ISF factor”, calculated following the flowcharts as explained in detail in section 03.).
- Below the autoISF section, the effective ISF (sens) results from dividing the (unchanged or adapted) ISF prior to “start autoISF”, with the determined “final ISF factor” at the end of the autoISF section of the SMB tab.

4.8 Profile helper

xls based tool is still under development / needs more user data / chapter will follow later